## Slide 1: ELRA
Dylan Ji
Endo Biliary RFA

## Slide 2
Q & A
From ELRA webinar with Dr. Reddy, Dr. Cho, Dr. Terry and Dr. Rungsan

## Slide 3: Q&A
Q. Do you place a plastic stent or metal stent after RFA ?
If you take a look at clinical articles and literatures, some of them placed metal stents and some of them placed only plastic stents. There is a variation. In our practice, we put a metal stent. But for hilar strictures, we put plastic stents repeatedly. The advantage of putting plastic stents is that you can repeat RfA after removing plastic stent and putting another stent. - Dr. Nageshwar Reddy-
In my cases, it depends on patient’s conditions. If patients are ok to take multiple ERCP, I placed plastic stents for multiple RFA. If patients are not good for advanced disease, I place a metal stent. - Dr. Jaehee Cho

## Slide 4: Q&A
Q. Generally, RFA can help to prevent the tumor ingrowth into metal stent so that it can increase stent patency. However, the cause of plastic stent blockage is different. It is usually caused by sludge, biofilm, and so on. So, How does RFA can prolong the stent patency of plastic stents?
-Dr. Reddy -
I do not know who plastic stent occlsion is affected by RFA. But RFA can maintain the longer patency of bile duct, so it may be related to the longer stent patency of plastic stents. We need to do more work on this. – Dr. Cho -
We ablate and put plastic stent for those patients who need multiple treatment of RFA. It is difficult to ablate whole lesion at one session. So, if need multiple ablation, we put plastic stent. - Dr. Rungsan -

## Slide 5: Q&A
Q. How many days do you do RFA after ampulectomy to ablate the residual tissue?
Our Korean group performed RFA about 2 weeks after ampullectomy.
Interestingly, we performed RFA inside pancreatic duct in some cases. In the cases, we reduced the ablation time and lowered the power. - Dr. Cho -
A. We also do RFA 2 weeks after ampullectomy. And when we do ampullary RFA, we always place plastic stent in the pancreatic duct. - Dr. Reddy -

## Slide 6: Q&A
Q. How long do you take an interval in multiple sessions in biliary stricture after plastic stents?
I take 2-3 months. But, after first session, the stent occlusion takes place sooner. So, I take 2 months after the first RFA. - Dr. Cho -

## Slide 7: Q&A
Q. Is it safe to do hilar RFA in patients with the tumor encasing portal vein and hepatic artery?
Hilar RFA, especially, for the extension to right intrahepatic duct, I am very careful of it. We do intraductal ultrasound to look at the anatomy of right hepatic artery before doing hilar RFA because the right duct is very close the artery. We had 2 major hemobilia cases with Habib and the both are from the right hepatic artery.
With portal vein, major portal vein never comes to the RFA field so it is not
major problem in this case. - Dr. Reddy -

## Slide 8: Q&A
Q. Have you ablated benign biliary strictures ?
I do not have any experience. - Dr. Cho -
No, either. There is one paper from China where biliary RFA was applied to benign strictures, especially post operative tight strictures. But the treatment is till not established and chance of perforation is very high. I do not recommend for benign strictures at present. - Dr. Nageshwar Reddy-

## Slide 9: Q&A
Q. Is it possible to have the RFA catheter while doing cholangioscope?
RFA is very blinded procedure. If the tissue does not contact the catheter, we don’t get actually good and effective RFA. So, if we can see, we can get very RFA. With the standard nasal cholangioscope, we can have RFA catheter inside. However, for spyglass scope, we don’t have any catheter. - Dr. Reddy –
We have some experience of percutaneous transhepatic cholangioscopy (PTC).
Along the scope, we insert the RFA catheter and do ablation under direct
visualization.

## Slide 10: Q&A
Q. After the ablation of ingrowth inside metal stent, why did you place a plastic stent rather than just leaving the metal stent?
The study (Kadayifci A et Endoscopy. 2016 Dec;48(12);1096-1101) showed good stent patency after RFA without placing plastic stent inside the metal stent. But we placed a plastic stent with a question whether or not there will be cholangitis afterwards. So the next step will be a compare plastic stenting vs no stenting in this circumstance cause the ablation is so good so we could leave the stent without plastic stent. - Dr. Terry Wong

## Slide 11: Q&A
Q. Is there any difference between ELRA and Habib catheter, especially for stent ingrowth?
I don’t have any experience with Habib probe. The advantage of ELRA is cut-off point. The RFA is cut off when you have very high impedance. When ELRA probe touch the metal stent, the impedance goes up very high and it is cut off. So you can be confident in terms of safety. - Dr. Terry Wong –
A. Habib catheter is 22mm only in size. Whereas, ELRA has four sizes. (11,18,22,33mm). RFA inside metal stent is difficult procedure because of electric shunt with metal stent. If you use short 11mm ELRA probe, you can reduce the shunt rate. -Dr. Cho -

## Slide 12: Q&A
Q. If you plan a RFA later on, is it better to use LCD stent?
Is it possible to do RFA going through the LCD stent?
LCD stent is good for the stent in stent in the management of hilar strictures. I have not put any ELRA probe through stent mesh of stent-in-stent LCD stent. But I wouldn’t suppose there will be any problem. - Dr. Terry Wong-
Q. Is there any technique to avoid the short circuit?
To use the short RFA probe and do multiple ablation is the best. The long probes have the higher chance of contacting the metal stent. - Dr. Terry Wong-

## Slide 13: Q&A
Q. Is there a comparation study to compare metal stenting in stent vs RFA in stent?
We have this data only where RFA group vs control plastic stent. We need to study with metal stent group. - Dr. Terry Wong –

## Slide 14: Q&A
Q. Stenting vs RFA + Stenting. The costs could be different. Is there any economic analysis on this?
We need outcome studies to determine whether or not RFA prolong the stent patency. This is the reason why I started to do RFA to answer the patients living 5 years, even 8 years with metal stent. And I run into problems with repeated stent occlusion, particularly, metal in metal. We need 3-10 years of stent patency. -Dr. Terry Wong-
We have increasing patients living 3 years, 4years, very common. And we bump into this problem of stent occlusion. So in fact, there is radical suggestion. Rather than putting a metal stent for patients who live long, why not be use of plastic stent and repeatedly change them and use RFA. That is a total change in view in this area, I think. -Dr. Reddy.

## Slide 15: Q&A
Q. What do you think of cytokine surge in biliary RFA?
We have to have a better idea of the mechanistic way of how RFA acts including local cytokine surge that we sometimes saw in the bile. Immunododulation seems to work in this in additional to the coagulation effect that is in the tissue. We need to study more on this. -Dr. Reddy-

## Slide 16
Thank you
for your attention!